# Supplementary Materials

Supplementary Table I. Buffer Matrix Ingredients

|  | Clinical study | Non-clinical studies |  |
| --- | --- | --- | --- |
| Component | Citrate saline buffer (mg/mL) | Citrate buffer without rHuPH20 (mg/mL) | Citrate buffer with rHuPH20 (mg/mL) |
| Sodium chloride | 5.8 | 8.77 | 8.77 |
| Polysorbate 80 | 0.4 | 0.3 | 0.3 |
| L-Arginine HCl | 10.5 | – | – |
| Sodium citrate dihydrate | 2.66 | 2.35 | 2.35 |
| Citric acid anhydrous | 0.17 | 0.385 | 0.385 |
| NaCMC | – | 10 | 10 |
| rHuPH20 | – | – | 2000 U/mL |
| Methionine | – | 1.49 | 1.49 |

HCl, hydrogen chloride; NaCMC, sodium carboxymethylcellulose; rHuPH20, recombinant human hyaluronidase PH20.

Supplementary Table II. Device Configurations in Miniature Pig Studies

| Device abbreviation | Description | Volume delivered (mL) | Injection depth (mm) | Needle bevel | Target injection duration (s)* |
| --- | --- | --- | --- | --- | --- |
| AI2 | Fully automated autoinjector | 2 | 5.5 | 27G | Slow (15 s) Fast (5 s) |
| sAI2/5.5 | Semi-automated autoinjector | 2 | 5.5 | 27G | Slow (15 s) Fast (5 s) |
| sAI2/7.5 | Semi-automated autoinjector | 2 | 7.5 | 27G | Slow (15 s) Fast (5 s) |
| AI1 | Fully automated autoinjector | 1 | 5.5 | 27G | 4−5 |
| PFS1 | Pre-filled syringe | 1 | 12.7† | 27G | 6 |
| PFS2 | Pre-filled syringe | 2 | 12.7† | 27G | 10 |

AI1, 1-mL auto-injector device; AI2, 2-mL auto-injector device; PFS1, 1-mL pre-filled syringe, PFS2, 2-mL pre-filled syringe; sAI2, semi-automated 2-mL auto-injector device.

*Original target injection durations, which are different from real performance data.

†PFS injection depth is not as precisely controlled as with AIs, as it involves a 45⁰ angle penetration with a 12.7-mm exposed length needle through pinched up skin, unlike vertical penetration of the AI needle without skin pinching.

Supplementary Table III. Device Configurations in Clinical Study

| Device abbreviation | Description | Volume delivered (mL) | Injection depth (mm) | Needle bevel | Target injection duration (s)* |
| --- | --- | --- | --- | --- | --- |
| AI2 fast | Fully automated autoinjector | 2 | 5.5 | 27G | 5 |
| hAI2 slow† | Fully automated autoinjector | 2 | 5.5 | 27G | 10 |
| PFS2 | Pre-filled syringe | 2 | 12.7‡ | 27G | 10 |

AI2, 2-mL auto-injector device; hAI2, human 2-mL auto-injector device; PFS2, 2-mL pre-filled syringe.

*Original target injection durations, which are different from real performance data.

†hAI2 slow use a stronger drive spring than the AI2 slow used in the miniature pig studies, hence a shorter injection duration (7 s target instead of 15 s).

‡PFS injection depth is not as precisely controlled as within AIs, as it involves a 45⁰ angle penetration with a 12.7-mm exposed length needle through pinched up skin, unlike vertical penetration of the AI needle without skin pinching up.

Supplementary Table IV. Grading Scales for Non-clinical Miniature Pig Studies for Erythema (a), Swelling Size (b), and Firmness (c)

a)

| Scale | Description |
| --- | --- |
| 0 | No erythema |
| 1 | Very slight erythema (barely perceptible) |
| 2 | Well-defined erythema |
| 3 | Moderate-to-severe erythema |
| 4 | Severe erythema (beet redness) to slight eschar formation |

b)

| Scale | Description |
| --- | --- |
| 0 | No swelling |
| 1 | Very slight swelling |
| 2 | Slight swelling |
| 3 | Moderate swelling |
| 4 | Severe swelling |

c)

| Scale | Description |
| --- | --- |
| 0 | No perceptible difference in firmness after injection |
| 1 | Very slightly firm (barely perceptible) |
| 2 | Mildly firm |
| 3 | Moderately firm |
| 4 | Very firm |

Supplementary Table V. Study Design Treatments for Clinical Evaluation of 
2-mL Subcutaneous Injections in Humans

| Treatment arm | Injections into the abdomen (sequence randomized) |
| --- | --- |
| 1 | 1 x PFS2 2 x AI2 fast 1 x hAI2 slow |
| 2 | 1 x PFS2 1 x AI2 fast 2 x hAI2 slow |

AI2, 2-mL auto-injector device; hAI2, human 2-mL auto-injector device; PFS2, 2-mL pre-filled syringe.

Duration between the injections was a minimum of 60 min.

Supplementary Fig. 1. Back leakage from 2-mL SC injections in humans by injection device

AI2, 2-mL auto-injector device; hAI2, human 2-mL auto-injector device; PFS2, 2-mL pre-filled syringe; 
SC, subcutaneous; SD, standard deviation.

Negative values are potentially due to liquid evaporation and are within measurement variability.